2016-7-21
12:30 pm
Bringing system up
Restart process I
Core B-37
21:05 Starting a new log file in 2016-07-21 folder. Mark Garnet fixed Mass Spec issue and I see significant Air in the Jacket so I am restarting the Jacket pump to purge the jacket system. REG

21:12 discovered that to get the recipe to connect to the HMI I had to reboot the RGA monitor. Now HMI is reporting the correct data. Also started purging the Jacket system. Opened up the rotometer needle valve to increase purge rate to >4LPM. REG

21:30 Set limits on jacket to 3 and 20PSI and adjusted the Ar Jacket makup regulator to allow the system to hit that limit in a few seconds. Loosened the connection between the MS and low pressure regulator to bleed air from Jacket system. REG

21:52 tightend the connection to the MS for a more accurate view of Jacket gas content.

22:04 changed input gas from H to He. REG
22:18 Turning on the chroma supplies REG

2016-07-22
10:30 started the MS cyclng between Core out and Jacket. REG
10:35 Opened bleed valve for Core Out just before the MS. REG
10:40 Closed bleed valve for core out. REG
14:23 Reduced Jacket in Pressure to stop purging the jacket. Reduced Jacket flow rate to 2 LPM. REG
16:10 Starting sequence OpenTerm000. REG

2016-07-28
Turned system back on. Sent mark email about not keeping purge on forever after alarm because drained Ar tank.

10:59 installed Jacket purge valve between equlabar regulator and the MS.
12:13 opened purge valves on Core and Jacket.

2016-08-04
10:00 Changed control setup to Minimum Wait After Sequencer Step = 15 min
Time for Stability = 40 min.

2016-08-05
13:20 Reduced Gain Schedule 1 on the temperature control from 2.4 to 2.3 W/degC.
17:55 Increased integration time on the Gain Schedule 1 on the temperature control from 11 to 15 minutes.
18:50 increase Gain Schedule 1 on the temperature control from 2.3 to 2.7 W/degC.

2016-08-06
13:40 Increased integration time on the Gain Schedule 1 on the temperature control from 15 to 20 minutes.

16:40 Decreased integration time on the Gain Schedule 1 on the temperature control from 20 to 10 minutes

2016-08-07
Changed over to new Ar tank. Noticed HMI was disconnected. On reconnection non of the flow meters was talking. Tried to reconnect the jacket and it completely stopped the jacket flow.