## Sheet1
A | Date & Time | 2016-09-07 21:31 | Value based on IPB2_Core_28b-_H2_150C-400C__Run1_day-01.csv row470 | Date & Time | 2016-09-05 09:01 | Value based on SRI 2016-08-12_day-24.csv row541
B | Box Hydrogen Level | 2.367 |  | Core Reactor Temp | 327.1972 | Temperature reported by thermocouple embedded in copper block that holds core. TC0 type K
C | Box Oxygen Level | 0.083302 |  | Core Gas Out | 47.908 | Temperature (degC) from thermocouple after the gas exits the core, and just before the heat exchanger. TC1.
D | Calorimeter Jacket Flowrate LPM | 6.1889 |  | Jacket Gas Out | 184.2295 | Temperature of gas entering the Argon cooling jacket in deg C, TC2 type N
E | Calorimeter Jacket H2O In T | 25.228 |  | Jacket Gas In | 31.777 | Temperature of gas entering the Argon cooling jacket in deg C. TC3, type K
F | Calorimeter Jacket H2O Out T | 25.3358 |  | Core Gas In | 51.8662 | Temperature (degC) from thermocouple after the core gas heater , and just before the tee into the core. TC4, k-type.
G | Calorimeter Jacket Power | -10.0258 | Calculated from the above three columns | Cooled Core Gas | 25.5488 | Temperature (deg C) measurement, thermocouple (TC5, n-type) located after the gas comes out of the core and through the heat exchanger
H | Core Q Power | 0 | (Vrms1 - Vrms2)* Vrms2 / Rterm | Inner Core Temp | 24.1973 | Temperature measured by thermocouple inserted into center of core tube. TC9, type K
I | Core Temp | 150.0122 | Temperature measured by thermocouple inserted into center of core tube. TC9, type K | TERMINATION XC H2O IN | 25.8249 | Temperature in degrees C of water entering termination heatsink. RTD0
J | Heater Manual Setpoint | 0 |  | TERMINATION XC H2O OUT | 26.8455 | temperature in degrees C of water entering termination heatsink. RTD1
K | Heater Power | 9.3358 | Power (Watts) used in the core gas heater. Chroma 61702 Ch2. | JCKT XC H2O IN | 25.6716 | Temperature of water entering the jacket gas to cooling water heat exchanger. RTD2
L | Heater Resistance | 0 |  | JCKT XC H2O OUT | 26.1399 | Temperature of water exiting the jacket gas to cooling water heat exchanger. Measured by an RTD3
M | Heater Watt-Hours | 0 |  | Core Pump Out Gas LPM | 0 | inapplicable in current stopped flow configuration
N | Inner Block Temp 1 | 141.5134 |  | Jckt Gas Circ LPM | 2.001 | Flowrate in Standard Liter/minute of Argon in the jacket gas loop
O | Inner Block Temp 2 | 141.6109 |  | Core Ar Makeup LPM | 0.0164 | Flow rate (Liters per minute) of Ar into the core. Sierra Air MFC, 0-1 LPM
P | Inner Outer Block Temp Diff | 112.7315 |  | H2 Makeup LPM | 0.02995 | Flow rate of Hydrogen or Helium in SLPM going into core
Q | Isoperibolic Calorimetry Power | 9.3959 | Calculated from the below three columns | Jckt Ar Makeup LPM | 0 | Flowrate in standard liter per minute of Argon entering Jacket cooling loop to make up gas lost to leaks
R | Outer Block Temp 1 | 28.8517 |  | Exit Gas LPM | 0 | When HHT was in circulating core flow configuration, it measured gas exiting core when the core circulation valve VL7 was not on
S | Outer Block Temp 2 | 28.8096 |  | Core Pump Inlet Press | 101.7122 | Inapplicable in current stopped flow configuration. Was measuring inlet pressure in PSIG of pump that circulated gas in core loop
T | Pressure Sensor PSI | 117.3697 |  | Core Out Press | 100.1025 | Pressure (in PSIG) measured after gas leaves the core, and the heat exchanger.
U | Pressure Setpoint | 0 |  | Core Out Press | 100.1025
V | Q Cur | 0.01159 | Current in Amps measured at Pi Filter | Jacket Out Press | 8.5228 | Pressure of gas exiting argon jacket in PSI
| Q kHz | 10 | Rep rate of q pulse in Kilohertz | Jacket In Press | 8.5523
| Q N Pulses | 1 |  | Core In Press | 100.7097 | Pressure (PSI) of gas going into the core. Measured before core gas heater. Analog input Ch5
| Q PCB Heatsink Flowrate LPM | 0.48514 |  | Core Pump Press | 59.4805 | In previous configuration with circulating core gas, measured pressure at exit of core circulation pump. Now just measures core pressure because it is in stopped flow configuration
| Q PCB Heatsink H2O In T | 25.2462 |  | Jckt HX Water LPM | 0.21778 | Flow of water in LPM going through heater exchanger that cools the Argon gas circulating through jacket. Most heat of jacket argon seems to be not lost in heat exchanger, but in tubing
| Q PCB Heatsink H2O Out T | 25.1874 |  | Termination HX Water LPM | 0.34385 | Flow rate in liter per minute of water flowing through termination heat sink. Heat dissipated by Q pulse in the termination
| Q Pow | 3.3978 | Q pulse Power in Watts measured at Pi Filter. The power that is going into the Q pulse board and termination | Core In H2 | 0 | Mass Spec measurement in mbar of H2 going into the core.
| Q Power Watt-Hours | 0 |  | Core In He | 0 | Mass Spec measurement in mbar of He going into the core.
| Q Pulse Delay (s) | 0 | if greater than 0, a number of pulses equal to “Q N pulses” will be fired, then no pulses will be fired for a time equal “Q Pulse Delay (s)” | Core In O | 0 | Mass Spec measurement in mbar of O going into the core.
| Q Pulse Length (ns) | 100 | length of Q pulse as generated at the CompactRIO digital output card, in nanoseconds | Core In H2O | 0 | Mass Spec measurement in mbar of H2O going into the core.
| Q Pulse PCB Heatsink Power | -2.2251 |  | Core In N | 0 | Mass Spec measurement in mbar of N going into the core
| Q Pulse Volt | 298.8857 | Voltage of Q pulse measured at the Pi Filter | Core In O2 | 0 | Mass Spec measurement in mbar of O2 going into the core.
| Q Set V | 300 |  | Core In Ar | 0 | Mass Spec measurement in mbar of Argon going into the core.
| Q Supply Power | 3.3864 |  | Core In M3 | 0 | Mass Spec measurement in mbar of ?? going into the core
| Q Supply Volt | 300.0064 | Q pulse voltage generated at the Chroma 61602 supply. This is AC voltage. | Core Out H2 | 0.0631 | Mass Spec measurement of H2 in gas leaving the core. Taken after the gas is cooled.
| Room Temperature | 25.3947 |  | Core Out He | 65.7 | Mass Spec measurement of He in gas leaving the core. Taken after the gas is cooled
| Temperature Controller Output | 0 |  | Core Out O | 0.0387 | Mass Spec measurement of O in gas leaving the core. Taken after the gas is cooled
| Temperature Controller PV | 150.0122 |  | Core Out H2O | 0.0659 | Mass Spec measurement of H2O in gas leaving the core. Taken after the gas is cooled
| Temperature Setpoint | 150 |  | Core Out N | 2.24 | Mass Spec measurement of N in gas leaving the core. Taken after the gas is cooled.
| Termination Heatsink Power | -0.25892 | Calculated from the below three columns | Core Out O2 | 0.437 | Mass Spec measurement of O2 in gas leaving the core. Taken after the gas is cooled.
| Termination Heatsink Flowrate LPM | 0.31038 |  | Core Out Ar | 0.0183 | Mass Spec measurement of Ar leaving the core. Taken after the gas is cooled.
| Termination Heatsink H2O In T | 25.2565 |  | Core Out M3 | 0.854 | Mass Spec measurement of M3 in gas leaving the core. Taken after the gas is cooled
| Termination Heatsink H2O Out T | 25.2632 |  | Jacket H2 | 0 | Mass Spectrometer reading of Hydrogen concentration in gas leaving the Argon cooling jacket
| Argon Purge Valve | 0 |  | Jacket He | 0 | Mass Spectrometer reading of Helium concentration in gas leaving the Argon cooling jacket
| Connected to Host | 1 |  | Jacket O | 0 | Mass Spectrometer reading of Monoatomic Oxygen concentration in gas leaving the Argon cooling jacket
| Core Vent Valve | 0 |  | Jacket H2O | 0 | Mass Spectrometer reading of H2O concentration in gas leaving the Argon cooling jacket
| Heater Enable | 1 |  | Jacket N | 0 | Mass Spectrometer reading of Nitrogen concentration in gas leaving the Argon cooling jacket
| Hydrogen Valves | 1 |  | Jacket O2 | 0 | Mass Spectrometer reading of Diatomic Oxygen concentration in gas leaving the Argon cooling jacket
| Power Controller Comm Status | 0 |  | Jacket Ar | 0 | Mass Spectrometer reading of Argon concentration in gas leaving the argon cooling jacket surrounding core
| Pressure Controller Comm Status | 0 |  | Jacket M3 | 0 | Mass Spectrometer reading of Mass 3 concentration in gas leaving the Argon cooling jacket
| Pressure Controller Mass Flow Overage | 0 |  | Cube H2 | 0 | Hydrogen concentration reported by MKS Mass Spectrometer when it is sampling from port 4 of rotary valve, which is connected to plexiglass box surrounding the reactor apparatus
| Pressure Controller Pressure Overage | 0 |  | Cube N | 0 | Nitrogen concentration reported by MKS Mass Spectrometer when it is sampling from port 4 of rotary valve, which is connected to plexiglass box surrounding the reactor apparatus
| Pressure Controller Temp Overage | 0 |  | Cube N | 0
| Q Enable | 0 |  | Room Temp | 25.1997 | Room temperature in degrees C
| Q Power Supply Comm Status | 1 |  | Pow In | 74.6936 | sum of all heater powers (all outputs of Chroma 61702) and Q power measured at Pi filter
| Temp Controller Comm Status | 1 |  | Pow Out | 31.5967 | Sum of Jacket Thermal Power and Termination Thermal Power/Totalized (integrated / summed) Argon passed through valve VA8, the jacket argon gas makeup valve
| Temperature Control Off-On | 1 |  | Jckt Ar Total | 78136.6893
| Temperature Controller Loop Break | 0 |  | Core Htr Pow | 41.5 | Power (Watts) used in the core gas heater. Chroma 61702 Ch2.
| Temperature Controller Over-Range | 0 |  | Jckt Htr Pow | 0
| Temperature Controller Remote SP Fail | 0 |  | Q Supply Volt | 99.99 | Q pulse voltage generated at the Chroma 61602 supply. This is AC voltage.
| Temperature Controller Sensor Break | 0 |  | Q kHz | 88.9548 | Rep rate of q pulse in Kilohertz
| Temperature Controller Alarm | 1 |  | Q N Pulses | 2 | Number of Q pulses fired between pauses if Q pulse Pulse Delay (s) is greater than 0. Does not matter if Q Pulse Delay (s) is 0
| Q Occurred? | 0 |  | Q Pulse Delay (s) | 0 | if greater than 0, a number of pulses equal to “Q N pulses” will be fired, then no pulses will be fired for a time equal “Q Pulse Delay (s)”
| Sequencing? | 1 |  | Q Pulse Length (ns) | 100 | length of Q pulse as generated at the CompactRIO digital output card, in nanoseconds
| Seq Step # | 2 |  | Jacket Case Temp | 1372 | temperature of jacket metal casing in degrees C. TC7, ype K
| Active Alarm Bits | 0 |  | Bias Volt | 0
|  |  |  | Bias Cur | 0
|  |  |  | Bias Pow | 0
|  |  |  | Core Htr Volt | 79.05 | Voltage (Volts) into the core heater. Chroma 61702 Ch1
|  |  |  | Core Htr Cur | 0.52 | Current (Amps) into the core heater. Chroma 61702 Ch1
|  |  |  | Core Htr Pow | 41.5 | Power (Watts) into the core heater. Chroma 61702 Ch1.
|  |  |  | Jckt Htr Volt | 0 | RMS Volts across jacket gas heater. Chroma 61702 Ch3
|  |  |  | Jckt Htr Cur | 0 | RMS Current in Amps going through jacket gas heater. Chroma 61702 Ch3.
|  |  |  | Jckt Htr Pow | 0 | Power in Watts dissipated in jacket gas heater. Chroma 61702 Ch3
|  |  |  | Q Cur | 0.16892 | Current in Amps measured at Pi Filter
|  |  |  | Q Pow | 32.7117 | Q pulse Power in Watts measured at Pi Filter. The power that is going into the Q pulse board and termination
|  |  |  | Q Pulse Volt | 126.2888 | Voltage of Q pulse measured at the Pi Filter
|  |  |  | Termination Therm Pow | 24.4815 | Cwater=4.186;% Specific heat of water => 1 calorie/gram °C = 4.186 joule/gram °C | Cwater=4.186;% Specific heat of water => 1 calorie/gram °C = 4.186 joule/gram °C
|  |  |  | 61602 Freq Hz | 240 |  | DT = TERMINATIONXCH2OOUT - TERMINATIONXCH2OIN; %C
|  |  |  | 61702 Freq Hz | 60 |  | Tflow = TerminationHXWaterLPM; %L/min = (1000/60) cc/sec = 0.0167 g/sec
|  |  |  | Q Supply Power | 40.2 |  | TTP = 4.184 * (Tflow*(1000/60) .* DT); %J/sec
|  |  |  | Press Ctrl Setpoint | 99.9957 | Setpoint in PSI of the electronic backpressure regulator regulating core pressure
|  |  |  | Press Ctrl Feedback | 100.0625 | Pressure in PSI measured by the electronic backpressure regulator
|  |  |  | Core Resistance | NaN
|  |  |  | Q Occurred? | 1 | Logical value (1/0) indicating whether a Q pulse has occurred within the last second
|  |  |  | ARGON GAS PURGE VALVE VA9 | 1 | Logical (1/0) for valve VA9 as indicated on system flow diagram
|  |  |  | ARGON GAS ROUTER VALVE VA6 | 1 | Logical (1/0) for valve VA6 as indicated on system flow diagram. Feeds Argon into core.
|  |  |  | ARGON GAS SOURCE VALVE VA5 | 1 | Logical (1/0) for valve VA5 as indicated on system flow diagram. In series with VA6. Feeds Argon into core
|  |  |  | ARGON JACKET FILL VALVE VA8 | 0 | Logical (1/0) for valve VA8. Feeds Argon into the jacket to maintain argon pressure. This valve is controlled by jacket by jacket pressure control loop
|  |  |  | ARGON PUMP EN | 1 | Logical (1/0) for Argon pump on/off
|  |  |  | BAND HEATER EN | 1 | Logical (1/0) for band heater on/off.
|  |  |  | CHROMA 61602 OUTPUT | 1 | Logical (1/0) for relay in series with Chroma 61602 power supply, which powers all Q pulse electronics
|  |  |  | BIAS EN | 0
|  |  |  | H GAS ROUTER VALVE VH4 | 1 | Logical state of valve VH4, lets Hydrogen or Helium into core. Should be renamed
|  |  |  | H GAS SOURCE VALVE VH3 | 0 | Logical state of valve VH3, lets Hydrogen into tee upstream of VH4.
|  |  |  | JACKET GAS HEATER EN | 0 | logical value (1/0) indicating status of relay wired into output of Chroma 61702 power supply phase 3. There is no longer a Jacket gas heater in use but channel can still be used for other purpose
|  |  |  | CORE GAS LOOP VL7 VALVE | 0
|  |  |  | WATER PUMP EN | 0 | Logical value (1/0) of relay controlling water pump circulating cooling water
|  |  |  | Active Alarm Bits | 0

## mark's note
this is the code I just wrote for HHT calorimetry. It's pretty simple, hopefully this is self-explanatory. Water heat capacity is assumed fixed at 4185.5 J/(kg*degC) and a density of 1 kg/liter. If this is ok with you all, I will roll it out in next update.
//Water calorimetry. Water heat capacity at 15C is 4185.5 J/(kg*degC)
//69.7583333 W/degC = 1 Liter/ minute * minute/(60 sec) * 1 kg/liter *4185.5 J/(kg*degC)
TermPower = 69.7583333*Term_LPM*(Term_Tout - Term_Tin) ;
JacketPower = 69.7583333*Jacket_LPM*(Jacket_HX_Tout-Jacket_HX_Tin);
// get water flow rate in kilograms per second. 1 Liter/minute * 1 kg/liter * 1 minute/(60 sec)
MassFlow_term= Term_LPM/60;
MassFlow_jacket= Jacket_LPM/60;
//Sum up all the electrical inputs. The 3 phase powers from the Chroma 61702, and Q power from the Pi filter
Power_in= Core_Htr_Pwr + Bias_V_Pwr + Jckt_Htr_Pwr + Q_pwr ;
//Power out is what comes out of the termination heat exchanger and the Jacket Argon to Water HeatXchanger
Power_out = TermPower + JacketPower ;
Excess_Power = Power_out - Power_in;