## IPB1-11022016
Date & Time | Format MM/DD/YYYY HH:mm:ss (e.g. 10/06/2015 07:30:18) Local time.
Box Hydrogen Level | Hydrogen level in % measured in containing IPB reactor. Not accurate due to difficulties with Argon being in box atmosphere. Ignore.
Box Oxygen Level | Relative percentage of Oxygen contained in the Argon gas that is purging the IPB's insulation space. Should be less than 1%. Measured by Apogee Instruments SO-120 electrolytic Oxygen sensor. Measured during Q pulse measurement interrupt period every 10 seconds
Calorimeter Jacket Flowrate LPM | Flow rate of water flowing through calorimeter outer jacket in Liters per minute. Measured during Q pulse measurement interrupt period every 10 seconds. Exponential moving average filtered (EMAF) with time constant of approximately 100s.
Calorimeter Jacket H2O In T | Temperature in degrees Celsius of water flowing into the calorimeter jacket. Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds.
Calorimeter Jacket H2O Out T | Temperature in degrees Celsius of water flowing out of the calorimeter jacket. Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds.
Calorimeter Jacket Power | Thermal power in Watts deposited by calorimeter into water jacket flow. Compensated for inlet and outlet RTD built-in temperature offset which is programmable though HMI. Will not be accurate because of large flowrate jacket as it is run right now. Programmable offset also needs to be current and updated. Only occasionally used as a sanity check.
Core Q Power | Calculated Joule heating caused by Q pulse heating in core. The RMS Voltage at the entrance and exit of the core in calculated from data acquired from an oscilloscope with two probes. Measured every 10 seconds. Core Q Power = (V1rms – V2rms) * V2rms / Zeffective Zeffective is calibrated such that the Coefficient of Performance is 1 when the core is running in Helium gas. Compensated for the Q pulse measurement off time to yield average Q pulse joule heating in core- Q pulse in on for 9.4s , off for 0.6s.Verification of this measurement is a work in progress. It correlates linearly with drop in heater power when running at constant temperature. More to come.
Core Q V1 rms | RMS voltage measured at entrance of the core Q pulse waveguide. Measured by oscilloscope
Core Q V2 rms | RMS voltage measured at exit of the core Q pulse waveguide. Measured by oscilloscope.
Core Temp | Temperature in degrees Celsius measured by type K Thermocouple at the center of the core. Thermocouple is built into core. Measured during Q pulse measurement interrupt period every 10 seconds
Heater Manual Setpoint | When running in constant power mode, this is the heater power setpoint in Watts.
Heater Power | Electrical power in Watts supplied by TDK-Lambda supply to the heater coil built into center of core. Derived from analog 4-20 mA voltage and current monitor signals out put by supply. Calibrated with manufacturer supplied power supply calibration data sheet. Averaged with a 1 second time constant EMAF.
Inner Block Temp 1 | Temperature in degrees Celsius measured by first type K Thermocouple embedded in core stainless steel or copper heater spreader. Measured during Q pulse measurement interrupt period every 10 seconds.
Inner Block Temp 2 | Temperature in degrees Celsius measured by second type K Thermocouple embedded in core heat stainless steel or copper spreader. Measured during Q pulse measurement interrupt period every 10 seconds. Right now it wired to same TC as first Inner Block TC, dues to mechanical modifications to reactor precluding presence of second TC. Will return in next mechanical iteration of reactor
Inner Outer Block Temp Diff | Average temperature difference in degrees celsius between calorimeter's inner core heat spreader and outer aluminum tube. Measured during Q pulse measurement interrupt period every 10 seconds. Inner Outer Block Temp Diff = (Inner Block Temp 1 + Inner Block Temp 2)/2 - (Outer Block Temp 1 + Outer Block Temp 2)/2, Used in calibration.
Isoperibolic Calorimetry Power | The heat flow out of the calorimeter in Watts calculated from the calibration polynomial. Isoperibolic Calorimetry Power = a0 + a1*x + a2*x^2 + a3*x^3 + a4*x^4, where x is the Inner-Outer Block Temp Diff.
Outer Block Temp 1 | Temperature in degrees Celsius measured by type K Thermocouple embedded in the calorimeter outer Aluminum block/tube. Measured during Q pulse measurement interrupt period every 10 seconds
Outer Block Temp 2 | Temperature in degrees Celsius measured by second type K Thermocouple embedded in the calorimeter outer Aluminum block/tube. Measured during Q pulse measurement interrupt period every 10 seconds
Pressure Sensor PSI | Pressure of gas inside core. In Pounds per Square Inch Gauge pressure.
Pressure Setpoint | Core pressure setpoint for systems with an electronic pressure controller. Otherwise pressure is regulated by manual regulator
Q Cur | DC current in Amps measured at input of Pi filter. Measured with a sense resistor.
Q kHz | Q pulse frequency in KiloHertz. Set by operator or by Q power control loop.
Q N Pulses | When running Q pulse in on/off mode, number of pulses to emit between off times. Mode rarely used but available.
Q PCB Heatsink Flowrate LPM | Flowrate of water flowing through the Q pulse printed circuit board transistor heat sinks in Liters per minute. Measured during Q pulse measurement interrupt period every 10 seconds. Exponential moving average filtered (EMAF) with time constant of approximately 100s
Q PCB Heatsink H2O In T | Temperature in degrees Celsius of water flowing into the Q pulse printed circuit board transistor heat sinks . Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds.
Q PCB Heatsink H2O Out T | Temperature in degrees Celsius of water flowing out of the Q pulse printed circuit board transistor heat sinks . Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds
Q Pow | Q pulse Power in Watts measured at Pi Filter. The power that is going into the Q pulse board and termination. Voltage and current signals are sampled at 10 Khz and averaged in 1000 sample blocks on the FPGA. The downsampled 10Hz power data is then piped to the real-time processor where it is averaged with an EMAF with a time constant of about 30 seconds. This averaged value is also used by the Q power control loop.
Q Pulse Delay (s) | If greater than 0, a number of pulses equal to “Q N pulses” will be fired, then no pulses will be fired for a time equal “Q Pulse Delay (s)”
Q Pulse Length (ns) | Length of Q pulse in nanoseconds generated at the CompactRIO NI 9401 digital output card.
Q Pulse PCB Heatsink Power | Thermal power in Watts deposited by the Q Pulse PCB heatsink water stream . Compensated for inlet and outlet RTD built-in temperature offset which is programmable though HMI.
Q Pulse Volt | Voltage of Q pulse measured at the Pi Filter.
Q Set V | Voltage setpoint of the Q DC power supply
Q Supply Power | Power output in Watts reported by the Q DC power supply. Supply has a serial interface. Queried every 10 seconds.
Q Supply Volt | Q pulse voltage generated at the Chroma 61602 supply. This is AC voltage.
Room Temperature | Ambient temperature in degrees Celsius measured by type K Thermocouple. Measured during Q pulse measurement interrupt period every 10 seconds
Temperature Controller Output | temperature controller output in percent. May be zero in some datafiles due to troubleshooting serial communications to controller.
Temperature Controller PV | Process Value the Eurotherm temperature controller is controlling off of. Same as Core temperature. Redundant.
Temperature Setpoint | Setpoint the Eurotherm temperature controls to. Controlled by operator or test sequencer
Termination Heatsink Power | Thermal power in Watts deposited by the termination heatsink water stream . Compensated for inlet and outlet RTD built-in temperature offset which is programmable though HMI.
Termination Heatsink Flowrate LPM | Flowrate of water flowing through the transmission line termination heat sink in Liters per minute. Measured during Q pulse measurement interrupt period every 10 seconds. Exponential moving average filtered (EMAF) with time constant of approximately 100s.
Termination Heatsink H2O In T | Temperature in degrees Celsius of water flowing into the transmission line termination heat sinks . Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds.
Termination Heatsink H2O Out T | Temperature in degrees Celsius of water flowing out of the transmission line termination heat sinks . Measured by a platinum resistive temperature device. Measured during Q pulse measurement interrupt period every 10 seconds.
Argon Purge Valve | Logical (1/0) for Argon valve to core. 1 is open 0 is closed.
Connected to Host | indicates whether cRIO is network connected to HMI program. For diagnostic purposes.
Core Vent Valve | Logical (1/0) for core vent valve. 1 is open 0 is closed. Open valve releases gas from core.
Heater Enable | Logical (1/0) for relay that controls current to heater. 1 in on 0 is off. Closed relay means no heater power enabled. Used for safety.
Hydrogen Valves | Logical state of valve VH3, lets Hydrogen into tee upstream of VH4.
Power Controller Comm Status | Logical (1/0) for Hydrogen valve to core. 1 is open 0 is closed.
Pressure Controller Comm Status
Pressure Controller Mass Flow Overage
Pressure Controller Pressure Overage
Pressure Controller Temp Overage
Q Enable | Logical (1/0) indicating whether FPGA is generating Q pulses
Q Power Supply Comm Status | Logical (1/0) indicating whether serial communications to Q pulse DC power supply is working or not. 1 is ok, 0 is serial communication error (timeout, parity error)
Temp Controller Comm Status | Logical (1/0) indicating whether serial communications to Eurotherm temperature controller is working or not. 1 is ok, 0 is serial communication error (timeout, parity error)
Temperature Control Off-On | Logical (1/0) indicating whether system is operating in temperature control mode or power control mode. 1 indicates temperature control mode, 0 power control mode. Temperature control mode heater power is controlled by Eurotherm controller to make core temperature equal to setpoint. In power control mode heater power supply is modulated to output constant power.
Temperature Controller Loop Break
Temperature Controller Over-Range | Eurotherm temperature controller diagnostics. Logical (1/0) , 1 is true
Temperature Controller Remote SP Fail
Temperature Controller Sensor Break
Temperature Controller Alarm
Q Occurred? | Logical (1/0) indicates whether a Q pulse was emitted since the last log entry. 1 is true.
Sequencing? | Logical (1/0) indicates whether the system is running a sequence file. 1 is true.
Seq Step Num | Step number of the current sequence being number
Active Alarm Bits